## Instructions
This application, DataLoader, requires user input on 2 worksheets only
Evaluationcritera
This worksheet is used to store the user defined values used in the assessment of both network and individual analyser acceptability.
The values needed are the test nmenoic used and level specific values of either uncertainty or relative uncertainty as appropriate.
These values need only be enterd once and will be re-used every time the application isused.
Updated values may need to be entered when different lot numbers of IQC are used.
The specific values entered are the reponsibility of the user based on their needs.
The values are essential if an automatic accept / reject interpretation is required
Remember - the values in the evaluation criteria sheet must match the levels being processed.
Config
This worksheet is the main control panel for this program
IQC data should be extracted from the users mid-ware / analysers before starting.
The config worksheet is split into 3 parts horizontally.
Column B contains instructions
Coulmn C requires data to be entered
Coulmn D contains support items such as buttons which facilitate file selection.
There are 4 main sections to this worksheet (vertically).
| 1 - a | Select the process type from a drop down list by clicking on the item (default = import and process) - this option is unlikely to be changed once set.
| b | Select the type of import file to be used. This is specific to a supplier / analyser. The correct selection is critical. - this option is unlikely to be changed once set.
| c | Enter the full name and path of the IQC data file extracted above. Click on the button to the right to obtain a file picker helper.
| d | Enter the full path where spreadsheets should be created when this application is run.
Files will be overwritten if files of the same name already exist in the destination folder.
Summary files will have a date/time stamp added when created.
It is strongly recommended that files are created in a shareable folder in the format x:\......\year\month where x = drive letter.
e.g. w:\monthly analyser comparison\2014\january. Keeping thois data will generally help IQC processes and support laboratory accreditation.
| e | Enter the full name and path of a blank copy of the Anova spreadsheet. This should be populated with the default coveraghe factors desired
A copy of this will be used for each analyte processed. - this option is unlikely to be changed once set.
| 2 | This section is used to enter the full path and name of any supplementary data files. Use of more than 1 file (entered in 1 - c) is optional
All files used must be of the same format. Test codes must be common and the same IQC material must be used per analyte.
This option allows multiple extracts to be combined and processed as a single entity.
Start at number 2 and work downwards without leaving gaps.
Click on a row to highlight it then click on the button to the right to obtain a file picker helper. This makes entry easier.
| 3 - a | This section allows entry of up to 5 QC lot number identifiers for which data will be eliminated.
The DataLoader will process only the first 2 levels of QC found per analyte by default. This option allows the user to effectively select the levels processed.
Using this may slow data processing but may be necessary if the analyser / mid-ware is unable to select levels to extract.
The QC identifier must be as shown in the importeddata sheet. It may be necessary to separate the import and process options if the values are not known at the start.
If the data extracts contain 3 or more levels of IQC data create 2 copies of the data loader and configure each to exclude the others data. Remember to use different output folders.
| b | Click on the number of SDs to obtaoin a drop down list. Select you preference. QC data outside the selected value will be eliminated from the process.
| 4 | Click on this button to start the process
Do not disturb the PC during the operation. A summary sheet will appear when preocessing is complete
Screen updating is disabled to prevent flickering / flashing
The current status of the program will be displayed in the status bar at the bottom of the Excel window
This application performs it's own housekeeping and data previously used is cleaned out at the start of processing.
Data being imported should not contain hidden or other characters which may be invalid for filenames.
Data files should not be opened anywhere else at the time of import as the Dataloader must have sole access
N.B if the program is interupted the macros will need to be reset. Close the workbook and start again.

## Config
Ver. 9.1 |  | Automated import and processing utility
| Instructions | Data entry | Support
1 - a | Select the processing option required | Import & process | (click on cell for dropdown option list)
b | Select your system supplier / option | Beckman (Datawizard) | (click on cell for dropdown option list)
c | Enter the filename and path of your QC data
(file paths must include drive letter and file extension)
d | Enter the full path to the location where
spreadsheets will be created
e | Enter the full path of the blank template file.
2 | Supplementary QC data files (optional) : | Leave this line blank
| Supplementary import file 2 |  | Use of more than 1 QC data file is optional. Data files must be of the same format. Enter the full path and filenames. Do not leave blank rows between files. Start at number 2. Click on the row required then click on the button below to select a file for that row.
Supplementary import file 3
Supplementary import file 4
Supplementary import file 5
Supplementary import file 6
Supplementary import file 7
Supplementary import file 8
Supplementary import file 9
Supplementary import file 10
Supplementary import file 11
Supplementary import file 12
3 - a | Delete data for the following QC Lot Nos / level (optional): Leave this line blank
| 1 |  | Enter the ID of the QC(s) to be deleted in lines labelled
| 2 |  | 1 - 5 opposite. Deleting data for specific levels / lot nos
| 3 |  | will slow down processing and is optional.
| 4 |  | Do not leave blank lines between entries
5
b | Tidy up your data: | Remove QC results greater or less than | 3 | SDs
4
Data import is limited to a maximum of 50000 results. Multiple analysers, tests and QC levels are processed simmultaneously.
Only the first 500 results for each of the first 2 levels of a QC material will be processed. All other values will be ignored.
A separate Anova file will be created for each test spanning multiple analysers and levels of QC in the export directory listed above.
Developed by Pete Ayling, Barking, Havering & Redbridge University Hospitals NHS Trust 2014

## evaluationcriteria
|  |  | Level 1 |  | Level 2 |  | Level 1 |  | Level 2
| User defined evaluation values | Test name | Within network uncertainty (AE27) | Relative within network uncertainty. (AE29) | Within network uncertainty (AF27) | Relative within network uncertainty. (AF29) | Acceptable within instr. uncertainty S.D. (I23) | Acceptable within instr. uncertainty C.V. (K23) | Acceptable within instr. uncertainty S.D. (T23) | Acceptable within instr. uncertainty C.V. (V23)
This sheet allows the user to define their own test specific values on
which the evaluation is based. These values need only be entered
once and will be re-used every time the application is used.
The application can create / update a list of test codes from a
previously run summary. This is done by clicking on the button
below and selecting a previous summary file. New test codes
will be added to the list each time this function is used - retaining
previously entered data.
Alternately, test names may be entered freehand - please ensure .
accuracy
The evaluation of the data will be done in real time during processing
using this data, if entered.
The values in brackets in the column headers refer to the cell in the
Anova spreadsheet where this data will be used.
These values may need to be updated when new lots of IQC are used.

## ImportData
Test | Analyser | Level 1 | Target | SD | Result